## Instructions
This spread sheet allows you to create a program to log 24 IR temperature sensor of 2 (12 in each) AM16/32 multiplexors connected to a CS1000 data logger
Programming Instructions
1 | Enter Plot Numbers into appropriate row in Sensor Coefficients tab
2 | Enter Scan frequency into the LoggerSettings tab
3 | Copy the cells hilighted Pink in the CRBasicTemplate tab
4 | Past these into Note pad and save as "FileName .CR1"
5 | Send this program to the datalogger using LoggerNet or equivelent communication software.
Wiring Instructions
Datalogger | Instrument | Instrument
Temp & Hum sensor
SE9 | Blue
SE10 | Red
G | Green
12V | Yellow
↓ | Clear
Radiation sensor
6H | Red
6L | Black
↓ | White
↓ | Clear
RainGauge
P1 | Blue
G | Grey
WindSpeed
P2 | Green
5V | Red
G | Clear
G | Blue
LeafWetness
SE15 | Red
VX1 | Black
G | purple
G | clear
Pyrradiometer
7H | Red
7L | Blue
SW12V | Pump positive
G | Pump negaitve
Datalogger | Multiplexor1 (4 x 16 mode)
C1 | clk
C2 | res
12V | 12V
G | G
1H | Com Odd H
1L | Com Odd L
3SE | Com Even H
VX2 | Com Even L
↓ | Com ↓ | Surface Temp Sensor
| 1H | red
| 1L | black
| 2H | green
| 2L | white
| ↓ | clear
| ↓ | blue
repeat above for each sensor for each 4X16 port set
Datalogger | Multiplexor2 (4 x 16 mode)
C1 | clk
C3 | res
12V | 12V
G | G
3H | Com Odd H
3L | Com Odd L
4SE | Com Even H
VX3 | Com Even L
↓ | Com ↓ | Surface Temp Sensor
| 1H | red
| 1L | black
| 2H | green
| 2L | white
| ↓ | clear
| ↓ | blue
repeat above for ports 1 - 8 (4X16 ports)
ports 9 -16 blank
Datalogger | Multiplexor3 (4 x 16 mode)
C1 | clk
C4 | res
12V | 12V
G | G
4H | Com Odd H
4L | Com Odd L
16SE | Com Even H
VX1 | Com Even L
↓ | Com ↓
Paraonometer
| 1H | Red | Signal
| 1L | Black | Signal Reference
| ↓ | White | Analog Ground
| ↓ | Clear | shield
repeat above for ports 1 - 12 (4X16 ports)
Temperature sensor
| 2H | Red | signal
| 2L | Black | voltage excitation
| ↓ | Purple | Analog Ground
| ↓ | Clear | shield
Only one of the even port will be used with this setup
Datalogger | Multiplexor4 (2 x 32 mode)
C1 | clk
C5 | res
12V | 12V
G | G
16H | Com Odd H
VX1 | Com Odd L
↓ | Com ↓
Temperature sensor
| 1H | Red | signal
| 1L | Black | voltage excitation
| ↓ | Purple | Analog Ground
| ↓ | Clear | shield
repeat above for ports 1 - 23 (2X32 ports)
Heat Flux Plate
| 25H | White | Signal
| 25L | -
| ↓ | Green | Signal Reference
| ↓ | Black | Shield
repeat above for ports 25 - 28 (2X32 ports)
Datalogger | Multiplexor5 (2 x 32 mode)
C1 | clk
C6 | res
12V | 12V
G | G
7H | Com Odd H
7L | Com Odd L
↓ | Com ↓
Linear Radiation Sensors
| 1H | Red | +
| 1L | Black | -
| ↓ | ?
repeat above for ports 1 - 27 (2X32 ports)
Thermocouple
| 28H | Purple | Signal
| 28L | Red | Signal Reference
| ↓ | Clear | Shield
repeat above for ports 28 - 31 (2X32 ports)

## Sensor coefficients
IR sensors | [5 empty] | 2 | 3 | 4 | 5 | 6 | 7 | 8
Array Ref | SensorRef | Multiplexor | Channel (Mux 4x16 channel that sensor is conected to) | Sensor Serial number | Plot ID | MC0 | MC1 | MC2 | BC0 | BC1 | BC2 | Cable length | Serial number | MC0 | MC1 | MC2 | BC0 | BC1 | BC2 | Cable length | Serial number | MC0 | MC1 | MC2 | BC0 | BC1 | BC2 | Cable length | Sensor Serial number | Plot ID | Mux | Channel
(1) | 1 | 1 | 1 | 3400 | 12 | 1336310000 | 7096880 | 75223.9 | -2761200 | -14756.1 | 14805.4 | 20 | 3379 | 1321320000 | 6942310 | 77904.1 | -9474180 | 85649.1 | 12783 | 30 | 3379 | 1321320000 | 6942310 | 77904.1 | -9474180 | 85649.1 | 12783 | 30 | 3400 | 12 | 1 | 1
(2) | 2 | 1 | 2 | 3401 | 17 | 1376890000 | 7239770 | 72395.2 | -5217560 | -45019.4 | 15545.3 | 20 | 3380 | 1395890000 | 7277220 | 75992 | -2662950 | -9950.63 | 13555.7 | 30 | 3380 | 1395890000 | 7277220 | 75992 | -2662950 | -9950.63 | 13555.7 | 30 | 3401 | 17 | 1 | 2
(3) | 3 | 1 | 3 | 3402 | 8 | 1392520000 | 7230830 | 75298.8 | -2209980 | -62020 | 14222.2 | 20 | 3381 | 1340100000 | 6959910 | 74279.9 | -683085 | 70789.2 | 13195.5 | 30 | 3381 | 1340100000 | 6959910 | 74279.9 | -683085 | 70789.2 | 13195.5 | 30 | 3402 | 8 | 1 | 3
(4) | 4 | 1 | 4 | 3399 | 7 | 1341860000 | 7045270 | 73401.4 | 647879 | 24049.6 | 14900.8 | 20 | 3382 | 1332630000 | 7031600 | 73620.1 | -1572190 | 37528.5 | 12417.8 | 30 | 3382 | 1332630000 | 7031600 | 73620.1 | -1572190 | 37528.5 | 12417.8 | 30 | 3399 | 7 | 1 | 4
(5) | 5 | 1 | 5 | 3385 | 1 | 1386130000 | 7242140 | 70955 | -2918440 | -11589.2 | 14149.3 | 30 | 3383 | 1414580000 | 7496330 | 79768.7 | -4488640 | 127752 | 14089.7 | 30 | 3383 | 1414580000 | 7496330 | 79768.7 | -4488640 | 127752 | 14089.7 | 30 | 3385 | 1 | 1 | 5
(6) | 6 | 1 | 6 | 3386 | 6 | 1368500000 | 7282460 | 75238.7 | -5148230 | 94378.8 | 14931.7 | 30 | 3384 | 1432430000 | 7402560 | 79259.3 | -6857680 | 133432 | 13799.7 | 30 | 3384 | 1432430000 | 7402560 | 79259.3 | -6857680 | 133432 | 13799.7 | 30 | 3386 | 6 | 1 | 6
(7) | 7 | 1 | 7 | 3388 | 5 | 1371730000 | 7058920 | 76906 | -6107980 | -35518.4 | 14153.7 | 30 | 3385 | 1386130000 | 7242140 | 70955 | -2918440 | -11589.2 | 14149.3 | 30 | 3385 | 1386130000 | 7242140 | 70955 | -2918440 | -11589.2 | 14149.3 | 30 | 3388 | 5 | 1 | 7
(8) | 8 | 1 | 8 | 3398 | 11 | 1434740000 | 7449330 | 73361.9 | -3427590 | 28290.2 | 15814.1 | 20 | 3386 | 1368500000 | 7282460 | 75238.7 | -5148230 | 94378.8 | 14931.7 | 30 | 3386 | 1368500000 | 7282460 | 75238.7 | -5148230 | 94378.8 | 14931.7 | 30 | 3398 | 11 | 1 | 8
(9) | 9 | 1 | 9 | 3396 | 18 | 1367620000 | 7237750 | 71645.9 | -5766930 | -21127.4 | 14807.4 | 20 | 3387 | 1321880000 | 6953780 | 75548.9 | -7079850 | 128729 | 13473.3 | 30 | 3387 | 1321880000 | 6953780 | 75548.9 | -7079850 | 128729 | 13473.3 | 30 | 3396 | 18 | 1 | 9
(10) | 10 | 1 | 10 | 3392 | 13 | 1339840000 | 7046660 | 74491.8 | -5004960 | 57560.6 | 13317.6 | 20 | 3388 | 1371730000 | 7058920 | 76906 | -6107980 | -35518.4 | 14153.7 | 30 | 3388 | 1371730000 | 7058920 | 76906 | -6107980 | -35518.4 | 14153.7 | 30 | 3392 | 13 | 1 | 10
(11) | 11 | 1 | 11 | 3394 | 14 | 1737340000 | 9126780 | 95247.2 | -18151800 | 888.145 | 15603.9 | 20 | 3389 | 1373910000 | 7204930 | 71235.4 | 2801350 | -42316.6 | 14731.4 | 30 | 3389 | 1373910000 | 7204930 | 71235.4 | 2801350 | -42316.6 | 14731.4 | 30 | 3394 | 14 | 1 | 11
(12) | 12 | 1 | 12 | 3393 | 10 | 1331080000 | 6996190 | 72127.4 | -8744260 | -40772.5 | 14361 | 20 | 3390 | 1410870000 | 7449540 | 77553.2 | -4211160 | 16325.6 | 12858.8 | 30 | 3390 | 1410870000 | 7449540 | 77553.2 | -4211160 | 16325.6 | 12858.8 | 30 | 3393 | 10 | 1 | 12
(13) | 13 | 1 | 13 | 3380 | 4 | 1395890000 | 7277220 | 75992 | -2662950 | -9950.63 | 13555.7 | 30 | 3391 | 1432990000 | 7521750 | 78210.1 | -3974070 | 3509.1 | 13660.6 | 20 | 3391 | 1432990000 | 7521750 | 78210.1 | -3974070 | 3509.1 | 13660.6 | 20 | 3380 | 4 | 1 | 13
(14) | 14 | 1 | 14 | 3389 | 2 | 1373910000 | 7204930 | 71235.4 | 2801350 | -42316.6 | 14731.4 | 30 | 3392 | 1339840000 | 7046660 | 74491.8 | -5004960 | 57560.6 | 13317.6 | 20 | 3392 | 1339840000 | 7046660 | 74491.8 | -5004960 | 57560.6 | 13317.6 | 20 | 3389 | 2 | 1 | 14
(15) | 15 | 1 | 15 | 3390 | 3 | 1410870000 | 7449540 | 77553.2 | -4211160 | 16325.6 | 12858.8 | 30 | 3393 | 1331080000 | 6996190 | 72127.4 | -8744260 | -40772.5 | 14361 | 20 | 3393 | 1331080000 | 6996190 | 72127.4 | -8744260 | -40772.5 | 14361 | 20 | 3390 | 3 | 1 | 15
(16) | 16 | 1 | 16 | 3387 | 22 | 1321880000 | 6953780 | 75548.9 | -7079850 | 128729 | 13473.3 | 30 | 3394 | 1737340000 | 9126780 | 95247.2 | -18151800 | 888.145 | 15603.9 | 20 | 3394 | 1737340000 | 9126780 | 95247.2 | -18151800 | 888.145 | 15603.9 | 20 | 3387 | 22 | 1 | 16
(17) | 17 | 2 | 1 | 3397 | 9 | 1347710000 | 7097850 | 73741.2 | -5754910 | -18618.3 | 13883.8 | 20 | 3395 | 1319160000 | 6990600 | 74259 | -3260600 | 3852.56 | 15757.4 | 20 | 3395 | 1319160000 | 6990600 | 74259 | -3260600 | 3852.56 | 15757.4 | 20 | 3397 | 9 | 2 | 1
(18) | 18 | 2 | 2 | 3381 | 23 | 1340100000 | 6959910 | 74279.9 | -683085 | 70789.2 | 13195.5 | 30 | 3396 | 1367620000 | 7237750 | 71645.9 | -5766930 | -21127.4 | 14807.4 | 20 | 3396 | 1367620000 | 7237750 | 71645.9 | -5766930 | -21127.4 | 14807.4 | 20 | 3381 | 23 | 2 | 2
(19) | 19 | 2 | 3 | 3379 | 19 | 1321320000 | 6942310 | 77904.1 | -9474180 | 85649.1 | 12783 | 30 | 3397 | 1347710000 | 7097850 | 73741.2 | -5754910 | -18618.3 | 13883.8 | 20 | 3397 | 1347710000 | 7097850 | 73741.2 | -5754910 | -18618.3 | 13883.8 | 20 | 3379 | 19 | 2 | 3
(20) | 20 | 2 | 4 | 3383 | 20 | 1414580000 | 7496330 | 79768.7 | -4488640 | 127752 | 14089.7 | 30 | 3398 | 1434740000 | 7449330 | 73361.9 | -3427590 | 28290.2 | 15814.1 | 20 | 3398 | 1434740000 | 7449330 | 73361.9 | -3427590 | 28290.2 | 15814.1 | 20 | 3383 | 20 | 2 | 4
(21) | 21 | 2 | 5 | 3396 | 15 | 1367620000 | 7237750 | 71645.9 | -5766930 | -21127.4 | 14807.4 | 20 | 3399 | 1341860000 | 7045270 | 73401.4 | 647879 | 24049.6 | 14900.8 | 20 | 3399 | 1341860000 | 7045270 | 73401.4 | 647879 | 24049.6 | 14900.8 | 20 | 3396 | 15 | 2 | 5
(22) | 22 | 2 | 6 | 3382 | 21 | 1332630000 | 7031600 | 73620.1 | -1572190 | 37528.5 | 12417.8 | 30 | 3400 | 1336310000 | 7096880 | 75223.9 | -2761200 | -14756.1 | 14805.4 | 20 | 3400 | 1336310000 | 7096880 | 75223.9 | -2761200 | -14756.1 | 14805.4 | 20 | 3382 | 21 | 2 | 6
(23) | 23 | 2 | 7 | 3391 | 16 | 1432990000 | 7521750 | 78210.1 | -3974070 | 3509.1 | 13660.6 | 20 | 3401 | 1376890000 | 7239770 | 72395.2 | -5217560 | -45019.4 | 15545.3 | 20 | 3401 | 1376890000 | 7239770 | 72395.2 | -5217560 | -45019.4 | 15545.3 | 20 | 3391 | 16 | 2 | 7
(24) | 24 | 2 | 8 | 3384 | 24 | 1432430000 | 7402560 | 79259.3 | -6857680 | 133432 | 13799.7 | 30 | 3402 | 1392520000 | 7230830 | 75298.8 | -2209980 | -62020 | 14222.2 | 20 | 3402 | 1392520000 | 7230830 | 75298.8 | -2209980 | -62020 | 14222.2 | 20 | 3384 | 24 | 2 | 8
Heat flux plates
| Sensor Ref | MultiPlex | Channel (Mux 4x16 channel that sensor is conected to) | Serial Number | Plot ID | 2 MC0: Multiplier |  |  | Calib coeff | [14 empty] | Serial Number | Plot ID | Mux | Channel
| 1 | 3 | 1 | 7677 | 19 | 2 MC0: 16.3881 |  | 7677 | 61.02 | [14 empty] | 7677 | 19 | 2 | 9
| 2 | 3 | 3 | 7678 | 20 | 2 MC0: 16.1005 |  | 7678 | 62.11 | [14 empty] | 7678 | 20 | 2 | 10
| 3 | 3 | 5 | 10135 | 17 | 2 MC0: 16.7084 |  | 7366 | 61.94 | [14 empty] | 10135 | 17 | 2 | 11
| 4 | 3 | 7 | 10136 | 13 | 2 MC0: 16.5344 |  | 7675 | 61.97 | [14 empty] | 10136 | 13 | 2 | 12
| 5 | 3 | 9 | 7364 | 5 | 2 MC0: 16.3881 |  | 7364 | 61.02 | [14 empty] | 7364 | 5 | 2 | 13
| 6 | 3 | 11 | 7363 | 6 | 2 MC0: 16.0798 |  | 7365 | 62.95 | [14 empty] | 7363 | 6 | 2 | 14
| 7 | 3 | 13 | 7675 | 2 | 2 MC0: 16.1368 |  | 7363 | 62.19 | [14 empty] | 7675 | 2 | 2 | 15
| 8 | 2 | 13 | 7366 | 4 | 2 MC0: 16.1447 |  | 7676 | 60.84 | [14 empty] | 7366 | 4 | 2 | 16
| 9 | 2 | 13 | 7676 | 22 | 2 MC0: 16.4366 |  | 10136 | 60.48 | [14 empty] | 7676 | 22 | 4 | 1
| 10 | 2 | 13 | 7365 | 21 | 2 MC0: 15.8856 |  | 10135 | 59.85 | [14 empty] | 7365 | 21 | 4 | 2
| 11 | 2 | 13 | 10138 | 15 | 2 MC0: 16.0591 |  | 10138 | 62.27 | [14 empty] | 10138 | 15 | 4 | 3
| 12 | 2 | 13 | 10137 | 10 | 2 MC0: 16.8719 |  | 10137 | 59.27 | [14 empty] | 10137 | 10 | 4 | 4
Pyrradiometer
1500
Note on E2: cflhxb: Enter the serial number of the sensor connected to each channel
Note on F2: cflhxb: Enter the Name of the plot to be displayed in the column head of the data field
Note on AJ2: cflhxb: Enter the serial number of the sensor connected to each channel
Note on AK2: cflhxb: Enter the Name of the plot to be displayed in the column head of the data field
Note on E29: cflhxb: Enter the serial number of the sensor connected to each channel
Note on F29: cflhxb: Enter the Name of the plot to be displayed in the column head of the data field
Note on AJ29: cflhxb: Enter the serial number of the sensor connected to each channel
Note on AK29: cflhxb: Enter the Name of the plot to be displayed in the column head of the data field

## LoggerSettings
ScanFrequenceValue | 2
ScanFrequenceUnits | min
DataAveragingFrequecy | 10 |  | 600
DataAveragingUnits | min
DataTableName | AllData
TempSignalSEChan | 9
HumSignalSEChan | 10
RadnDiffChan | 6
RainMmPerTip | 1
RainSignalPulseChan | 1
WindSignalPulseChan | 2
WindMSPerPulse | 0.765
WindOffset | 0.35
LeafWetnessSEChan | 15
LeafWetnessVXChan | 1
ClockComChannel | 1
PyrradiometerDiffChan | 7
MultiPlexor | 1 | 2 | 3 | 4 | 5
MUXResComChannel | 2 | 3 | 4 | 5 | 6
MUXDiffChannel | 1 | 3 | 4 | None | 7
MUXSEChannel | 3 | 4 | 16 | 16 | None
MUXVXChannel | 2 | 3 | 1 | 1 | None
Note on A4: cflhxb: Can be mSec, Sec or Min
Note on A6: cflhxb: This must be a multiple of the Scan frequency
Note on A8: cflhxb: Can be mSec, Sec or Min

## CRBasicTemplate
'CR1000
'Created by Hamish Brown using a template made in Excel
'Declare Variables and Units
Dim TT_K(24)
Dim SBT_K(24)
Dim m(24)
Dim b(24)
Public BattV
Public PTemp_C
Public AirTC
Public RH
Public SlrkW
Public SlrMJ
Public Rain_mm
Public WS_ms
Public LeafWet_kohms
Public NetRadn
Public TT_C(24)
Public SBT_C(24)
Public TTmV(24)
Public SoilAvTC(12)
Public HeatFlux(12)
Public SlrReflectkW(12)
Public SlrReflectMJ(12)
Public SoilSurfTC(24)
Public LinPAR(27)
Units BattV=Volts
Units PTemp_C=Deg C
Units AirTC=Deg C
Units RH=%
Units SlrkW=kW/m^2
Units SlrMJ=MJ/m^2
Units Rain_mm=mm
Units WS_ms=meters/second
Units LeafWet_kohms=kilohms
Units NetRadn=W/m^2
Units TT_C=Deg C
Units SBT_C=Deg C
Units TTmV=mV
Units SoilAvTC=DegC
Units HeatFlux=W/m^2
Units SlrReflectkW=kW/m^2
Units SlrReflectMJ=MJ/m^2
Units SoilSurfTC=Deg C
Units LinPAR=mmolPAR
'Define Data Tables
DataTable(AllData,True,-1)
DataInterval(0,10,min,10)
Average(1,BattV,FP2,False)
Average(1,PTemp_C,FP2,False)
Average(1,AirTC,FP2,False)
Maximum(1,RH,FP2,False,False)
Average(1,SlrkW,FP2,False)
Totalize(1,SlrMJ,IEEE4,False)
Totalize(1,Rain_mm,FP2,False)
Average(1,WS_ms,FP2,False)
Average(1,LeafWet_kohms,FP2,False)
Average(1,NetRadn,FP2,False)
| Average(1,SBT_C( | 1 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(1),FP2,False)
| Average(1,SBT_C( | 2 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(2),FP2,False)
| Average(1,SBT_C( | 3 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(3),FP2,False)
| Average(1,SBT_C( | 4 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(4),FP2,False)
| Average(1,SBT_C( | 5 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(5),FP2,False)
| Average(1,SBT_C( | 6 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(6),FP2,False)
| Average(1,SBT_C( | 7 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(7),FP2,False)
| Average(1,SBT_C( | 8 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(8),FP2,False)
| Average(1,SBT_C( | 9 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(9),FP2,False)
| Average(1,SBT_C( | 10 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(10),FP2,False)
| Average(1,SBT_C( | 11 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(11),FP2,False)
| Average(1,SBT_C( | 12 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(12),FP2,False)
| Average(1,SBT_C( | 13 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(13),FP2,False)
| Average(1,SBT_C( | 14 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(14),FP2,False)
| Average(1,SBT_C( | 15 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(15),FP2,False)
| Average(1,SBT_C( | 16 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(16),FP2,False)
| Average(1,SBT_C( | 17 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(17),FP2,False)
| Average(1,SBT_C( | 18 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(18),FP2,False)
| Average(1,SBT_C( | 19 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(19),FP2,False)
| Average(1,SBT_C( | 20 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(20),FP2,False)
| Average(1,SBT_C( | 21 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(21),FP2,False)
| Average(1,SBT_C( | 22 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(22),FP2,False)
| Average(1,SBT_C( | 23 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(23),FP2,False)
| Average(1,SBT_C( | 24 |  | ),FP2,False) |  |  |  | Average(1,SBT_C(24),FP2,False)
| Average(1,TT_C( | 1 |  | ),FP2,False) |  |  |  | Average(1,TT_C(1),FP2,False)
| Average(1,TT_C( | 2 |  | ),FP2,False) |  |  |  | Average(1,TT_C(2),FP2,False)
| Average(1,TT_C( | 3 |  | ),FP2,False) |  |  |  | Average(1,TT_C(3),FP2,False)
| Average(1,TT_C( | 4 |  | ),FP2,False) |  |  |  | Average(1,TT_C(4),FP2,False)
| Average(1,TT_C( | 5 |  | ),FP2,False) |  |  |  | Average(1,TT_C(5),FP2,False)
| Average(1,TT_C( | 6 |  | ),FP2,False) |  |  |  | Average(1,TT_C(6),FP2,False)
| Average(1,TT_C( | 7 |  | ),FP2,False) |  |  |  | Average(1,TT_C(7),FP2,False)
| Average(1,TT_C( | 8 |  | ),FP2,False) |  |  |  | Average(1,TT_C(8),FP2,False)
| Average(1,TT_C( | 9 |  | ),FP2,False) |  |  |  | Average(1,TT_C(9),FP2,False)
| Average(1,TT_C( | 10 |  | ),FP2,False) |  |  |  | Average(1,TT_C(10),FP2,False)
| Average(1,TT_C( | 11 |  | ),FP2,False) |  |  |  | Average(1,TT_C(11),FP2,False)
| Average(1,TT_C( | 12 |  | ),FP2,False) |  |  |  | Average(1,TT_C(12),FP2,False)
| Average(1,TT_C( | 13 |  | ),FP2,False) |  |  |  | Average(1,TT_C(13),FP2,False)
| Average(1,TT_C( | 14 |  | ),FP2,False) |  |  |  | Average(1,TT_C(14),FP2,False)
| Average(1,TT_C( | 15 |  | ),FP2,False) |  |  |  | Average(1,TT_C(15),FP2,False)
| Average(1,TT_C( | 16 |  | ),FP2,False) |  |  |  | Average(1,TT_C(16),FP2,False)
| Average(1,TT_C( | 17 |  | ),FP2,False) |  |  |  | Average(1,TT_C(17),FP2,False)
| Average(1,TT_C( | 18 |  | ),FP2,False) |  |  |  | Average(1,TT_C(18),FP2,False)
| Average(1,TT_C( | 19 |  | ),FP2,False) |  |  |  | Average(1,TT_C(19),FP2,False)
| Average(1,TT_C( | 20 |  | ),FP2,False) |  |  |  | Average(1,TT_C(20),FP2,False)
| Average(1,TT_C( | 21 |  | ),FP2,False) |  |  |  | Average(1,TT_C(21),FP2,False)
| Average(1,TT_C( | 22 |  | ),FP2,False) |  |  |  | Average(1,TT_C(22),FP2,False)
| Average(1,TT_C( | 23 |  | ),FP2,False) |  |  |  | Average(1,TT_C(23),FP2,False)
| Average(1,TT_C( | 24 |  | ),FP2,False) |  |  |  | Average(1,TT_C(24),FP2,False)
| Average(1,TTmV( | 1 |  | ),FP2,False) |  |  |  | Average(1,TTmV(1),FP2,False)
| Average(1,TTmV( | 2 |  | ),FP2,False) |  |  |  | Average(1,TTmV(2),FP2,False)
| Average(1,TTmV( | 3 |  | ),FP2,False) |  |  |  | Average(1,TTmV(3),FP2,False)
| Average(1,TTmV( | 4 |  | ),FP2,False) |  |  |  | Average(1,TTmV(4),FP2,False)
| Average(1,TTmV( | 5 |  | ),FP2,False) |  |  |  | Average(1,TTmV(5),FP2,False)
| Average(1,TTmV( | 6 |  | ),FP2,False) |  |  |  | Average(1,TTmV(6),FP2,False)
| Average(1,TTmV( | 7 |  | ),FP2,False) |  |  |  | Average(1,TTmV(7),FP2,False)
| Average(1,TTmV( | 8 |  | ),FP2,False) |  |  |  | Average(1,TTmV(8),FP2,False)
| Average(1,TTmV( | 9 |  | ),FP2,False) |  |  |  | Average(1,TTmV(9),FP2,False)
| Average(1,TTmV( | 10 |  | ),FP2,False) |  |  |  | Average(1,TTmV(10),FP2,False)
| Average(1,TTmV( | 11 |  | ),FP2,False) |  |  |  | Average(1,TTmV(11),FP2,False)
| Average(1,TTmV( | 12 |  | ),FP2,False) |  |  |  | Average(1,TTmV(12),FP2,False)
| Average(1,TTmV( | 13 |  | ),FP2,False) |  |  |  | Average(1,TTmV(13),FP2,False)
| Average(1,TTmV( | 14 |  | ),FP2,False) |  |  |  | Average(1,TTmV(14),FP2,False)
| Average(1,TTmV( | 15 |  | ),FP2,False) |  |  |  | Average(1,TTmV(15),FP2,False)
| Average(1,TTmV( | 16 |  | ),FP2,False) |  |  |  | Average(1,TTmV(16),FP2,False)
| Average(1,TTmV( | 17 |  | ),FP2,False) |  |  |  | Average(1,TTmV(17),FP2,False)
| Average(1,TTmV( | 18 |  | ),FP2,False) |  |  |  | Average(1,TTmV(18),FP2,False)
| Average(1,TTmV( | 19 |  | ),FP2,False) |  |  |  | Average(1,TTmV(19),FP2,False)
| Average(1,TTmV( | 20 |  | ),FP2,False) |  |  |  | Average(1,TTmV(20),FP2,False)
| Average(1,TTmV( | 21 |  | ),FP2,False) |  |  |  | Average(1,TTmV(21),FP2,False)
| Average(1,TTmV( | 22 |  | ),FP2,False) |  |  |  | Average(1,TTmV(22),FP2,False)
| Average(1,TTmV( | 23 |  | ),FP2,False) |  |  |  | Average(1,TTmV(23),FP2,False)
| Average(1,TTmV( | 24 |  | ),FP2,False) |  |  |  | Average(1,TTmV(24),FP2,False)
| Average(1, SoilSurfTC( | 1 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(1),FP2,False)
| Average(1, SoilSurfTC( | 2 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(2),FP2,False)
| Average(1, SoilSurfTC( | 3 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(3),FP2,False)
| Average(1, SoilSurfTC( | 4 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(4),FP2,False)
| Average(1, SoilSurfTC( | 5 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(5),FP2,False)
| Average(1, SoilSurfTC( | 6 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(6),FP2,False)
| Average(1, SoilSurfTC( | 7 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(7),FP2,False)
| Average(1, SoilSurfTC( | 8 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(8),FP2,False)
| Average(1, SoilSurfTC( | 9 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(9),FP2,False)
| Average(1, SoilSurfTC( | 10 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(10),FP2,False)
| Average(1, SoilSurfTC( | 11 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(11),FP2,False)
| Average(1, SoilSurfTC( | 12 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(12),FP2,False)
| Average(1, SoilSurfTC( | 13 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(13),FP2,False)
| Average(1, SoilSurfTC( | 14 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(14),FP2,False)
| Average(1, SoilSurfTC( | 15 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(15),FP2,False)
| Average(1, SoilSurfTC( | 16 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(16),FP2,False)
| Average(1, SoilSurfTC( | 17 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(17),FP2,False)
| Average(1, SoilSurfTC( | 18 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(18),FP2,False)
| Average(1, SoilSurfTC( | 19 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(19),FP2,False)
| Average(1, SoilSurfTC( | 20 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(20),FP2,False)
| Average(1, SoilSurfTC( | 21 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(21),FP2,False)
| Average(1, SoilSurfTC( | 22 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(22),FP2,False)
| Average(1, SoilSurfTC( | 23 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(23),FP2,False)
| Average(1, SoilSurfTC( | 24 |  | ),FP2,False) |  |  |  | Average(1, SoilSurfTC(24),FP2,False)
| Average(1,HeatFlux( | 1 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(1),FP2,False)
| Average(1,HeatFlux( | 2 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(2),FP2,False)
| Average(1,HeatFlux( | 3 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(3),FP2,False)
| Average(1,HeatFlux( | 4 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(4),FP2,False)
| Average(1,HeatFlux( | 5 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(5),FP2,False)
| Average(1,HeatFlux( | 6 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(6),FP2,False)
| Average(1,HeatFlux( | 7 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(7),FP2,False)
| Average(1,HeatFlux( | 8 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(8),FP2,False)
| Average(1,HeatFlux( | 9 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(9),FP2,False)
| Average(1,HeatFlux( | 10 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(10),FP2,False)
| Average(1,HeatFlux( | 11 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(11),FP2,False)
| Average(1,HeatFlux( | 12 |  | ),FP2,False) |  |  |  | Average(1,HeatFlux(12),FP2,False)
| Average(1,SlrReflectkW( | 1 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(1),FP2,False)
| Average(1,SlrReflectkW( | 2 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(2),FP2,False)
| Average(1,SlrReflectkW( | 3 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(3),FP2,False)
| Average(1,SlrReflectkW( | 4 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(4),FP2,False)
| Average(1,SlrReflectkW( | 5 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(5),FP2,False)
| Average(1,SlrReflectkW( | 6 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(6),FP2,False)
| Average(1,SlrReflectkW( | 7 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(7),FP2,False)
| Average(1,SlrReflectkW( | 8 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(8),FP2,False)
| Average(1,SlrReflectkW( | 9 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(9),FP2,False)
| Average(1,SlrReflectkW( | 10 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(10),FP2,False)
| Average(1,SlrReflectkW( | 11 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(11),FP2,False)
| Average(1,SlrReflectkW( | 12 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectkW(12),FP2,False)
| Average(1,SlrReflectMJ( | 1 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(1),FP2,False)
| Average(1,SlrReflectMJ( | 2 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(2),FP2,False)
| Average(1,SlrReflectMJ( | 3 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(3),FP2,False)
| Average(1,SlrReflectMJ( | 4 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(4),FP2,False)
| Average(1,SlrReflectMJ( | 5 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(5),FP2,False)
| Average(1,SlrReflectMJ( | 6 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(6),FP2,False)
| Average(1,SlrReflectMJ( | 7 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(7),FP2,False)
| Average(1,SlrReflectMJ( | 8 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(8),FP2,False)
| Average(1,SlrReflectMJ( | 9 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(9),FP2,False)
| Average(1,SlrReflectMJ( | 10 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(10),FP2,False)
| Average(1,SlrReflectMJ( | 11 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(11),FP2,False)
| Average(1,SlrReflectMJ( | 12 |  | ),FP2,False) |  |  |  | Average(1,SlrReflectMJ(12),FP2,False)
| Average(1,SoilAvTC( | 1 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(1),FP2,False)
| Average(1,SoilAvTC( | 2 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(2),FP2,False)
| Average(1,SoilAvTC( | 3 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(3),FP2,False)
| Average(1,SoilAvTC( | 4 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(4),FP2,False)
| Average(1,SoilAvTC( | 5 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(5),FP2,False)
| Average(1,SoilAvTC( | 6 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(6),FP2,False)
| Average(1,SoilAvTC( | 7 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(7),FP2,False)
| Average(1,SoilAvTC( | 8 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(8),FP2,False)
| Average(1,SoilAvTC( | 9 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(9),FP2,False)
| Average(1,SoilAvTC( | 10 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(10),FP2,False)
| Average(1,SoilAvTC( | 11 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(11),FP2,False)
| Average(1,SoilAvTC( | 12 |  | ),FP2,False) |  |  |  | Average(1,SoilAvTC(12),FP2,False)
| Average(1,LinPAR( | 1 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(1),FP2,False)
| Average(1,LinPAR( | 2 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(2),FP2,False)
| Average(1,LinPAR( | 3 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(3),FP2,False)
| Average(1,LinPAR( | 4 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(4),FP2,False)
| Average(1,LinPAR( | 5 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(5),FP2,False)
| Average(1,LinPAR( | 6 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(6),FP2,False)
| Average(1,LinPAR( | 7 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(7),FP2,False)
| Average(1,LinPAR( | 8 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(8),FP2,False)
| Average(1,LinPAR( | 9 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(9),FP2,False)
| Average(1,LinPAR( | 10 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(10),FP2,False)
| Average(1,LinPAR( | 11 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(11),FP2,False)
| Average(1,LinPAR( | 12 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(12),FP2,False)
| Average(1,LinPAR( | 13 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(13),FP2,False)
| Average(1,LinPAR( | 14 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(14),FP2,False)
| Average(1,LinPAR( | 15 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(15),FP2,False)
| Average(1,LinPAR( | 16 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(16),FP2,False)
| Average(1,LinPAR( | 17 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(17),FP2,False)
| Average(1,LinPAR( | 18 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(18),FP2,False)
| Average(1,LinPAR( | 19 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(19),FP2,False)
| Average(1,LinPAR( | 20 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(20),FP2,False)
| Average(1,LinPAR( | 21 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(21),FP2,False)
| Average(1,LinPAR( | 22 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(22),FP2,False)
| Average(1,LinPAR( | 23 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(23),FP2,False)
| Average(1,LinPAR( | 24 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(24),FP2,False)
| Average(1,LinPAR( | 25 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(25),FP2,False)
| Average(1,LinPAR( | 26 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(26),FP2,False)
| Average(1,LinPAR( | 27 |  | ),FP2,False) |  |  |  | Average(1,LinPAR(27),FP2,False)
EndTable
'Main Program
BeginProg
'Main Scan. Scan(Internal, Units, BufferSize, Count)
Scan( | 2 | , | min | ,1,0) |  |  |  | Scan(2,min,1,0)
'Default Datalogger Battery Voltage measurement 'BattV'
Battery(BattV)
'Default Wiring Panel Temperature measurement 'PTemp_C'
PanelTemp(PTemp_C,_60Hz)
'CS500 Temperature & Relative Humidity Sensor measurements 'AirTC' and 'RH'. VoltSe(Dest, Reps, Range, SEChan, MeasOffset, SettlingTime, Integ, Mult, Offset)
VoltSE(AirTC,1,mV2500, | 9 | ,0,0,_60Hz,0.1,-40) | [6 empty] | VoltSE(AirTC,1,mV2500,9,0,0,_60Hz,0.1,-40)
VoltSE(RH,1,mV2500, | 10 | ,0,0,_60Hz,0.1,0) | [6 empty] | VoltSE(RH,1,mV2500,10,0,0,_60Hz,0.1,0)
If (RH>100) And (RH<108) Then RH=100
'LI200X Pyranometer measurements 'SlrMJ' and 'SlrkW'. Voltdiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Ofset)
VoltDiff(SlrkW,1,mV7_5, | 6 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrkW,1,mV7_5,6,True,0,_60Hz,1,0)
| 0.0002 | [7 empty] | If SlrkW<0 Then SlrkW=0
SlrMJ=SlrkW* | 0.12 | [7 empty] | SlrMJ=SlrkW*0.12
SlrkW=SlrkW*0.2
'Generic Tipping Bucket Rain Gauge measurement 'Rain_mm'. PulseCount(Dest, Reps, Pchan, Pconfig, Poption, Mult, Offset)
PulseCount(Rain_mm,1, | 1 | ,2,0, | 1 | ,0) |  |  |  |  | PulseCount(Rain_mm,1,1,2,0,1,0)
'NRG #40 Wind Speed Sensor measurements 'WS_ms'. PulseCount(Dest, Reps, Pchan, Pconfig, Poption, Mult, Offset)
PulseCount(WS_ms,1, | 2 | ,2,1, | 0.765 | , | 0.35 | ) |  |  | PulseCount(WS_ms,1,2,2,1,0.765,0.35)
If WS_ms<0.36 Then WS_ms=0
'237 Leaf Wetness Sensor measurement 'LeafWet_kohms'. BrHalf(Dest, Reps, Range, SEChan, VX/ExChan, MeasPEx, ExmV, RevEx, Settlingtime, Integ, Mult, Offset)
BrHalf(LeafWet_kohms,1,mV25, | 15 | , | 1 | ,1,2500,True,0,250,1,0) |  |  |  |  | BrHalf(LeafWet_kohms,1,mV25,15,1,1,2500,True,0,250,1,0)
LeafWet_kohms=(1/LeafWet_kohms)-101
'CN1-R net Pyrradiometer. Voltdiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Ofset)
VoltDiff(NetRadn,1,mv250, | 7 | ,True,0,_60Hz, | 1500 | ,0) |  |  |  |  | VoltDiff(NetRadn,1,mv250,7,True,0,_60Hz,1500,0)
MUXNumber | 1 | IR sensors | [6 empty] | 'Turn AM16/32 Multiplexer On. PortSet(DataLogger com channel, State 1=on and can respond to clock pulses)
PortSet( | 2 | ,1) | [6 empty] | PortSet(2,1)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 1 | 1 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 1 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(1),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 1 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(1),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(1)= | 1336310000 | +( | 7096880 | *SBT_C(1))+( | 75223.9 | *SBT_C(1)^2) |  |  |  |  | m(1)=1336310000+(7096880*SBT_C(1))+(75223.9*SBT_C(1)^2)
b(1)= | -2761200 | +( | -14756.1 | *SBT_C(1))+( | 14805.4 | *SBT_C(1)^2) |  |  |  |  | b(1)=-2761200+(-14756.1*SBT_C(1))+(14805.4*SBT_C(1)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 1 | )=SBT_C( | 1 | )+273.15 | [6 empty] | SBT_K(1)=SBT_C(1)+273.15
TT_K( | 1 | )=SBT_K( | 1 | )^4+TTmV(1) | *m(1) | +b(1) |  |  |  |  | TT_K(1)=SBT_K(1)^4+TTmV(1)*m(1)+b(1)
TT_K( | 1 | )=SQR(SQR(TT_K( | 1 | ))) | [6 empty] | TT_K(1)=SQR(SQR(TT_K(1)))
'Convert target temperature into desired units
TT_C( | 1 | )=TT_K( | 1 | )-273.15 | [6 empty] | TT_C(1)=TT_K(1)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 2 | 2 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 2 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(2),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 2 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(2),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(2)= | 1376890000 | +( | 7239770 | *SBT_C(2))+( | 72395.2 | *SBT_C(2)^2) |  |  |  |  | m(2)=1376890000+(7239770*SBT_C(2))+(72395.2*SBT_C(2)^2)
b(2)= | -5217560 | +( | -45019.4 | *SBT_C(2))+( | 15545.3 | *SBT_C(2)^2) |  |  |  |  | b(2)=-5217560+(-45019.4*SBT_C(2))+(15545.3*SBT_C(2)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 2 | )=SBT_C( | 2 | )+273.15 | [6 empty] | SBT_K(2)=SBT_C(2)+273.15
TT_K( | 2 | )=SBT_K( | 2 | )^4+TTmV(2) | *m(2) | +b(2) |  |  |  |  | TT_K(2)=SBT_K(2)^4+TTmV(2)*m(2)+b(2)
TT_K( | 2 | )=SQR(SQR(TT_K( | 2 | ))) | [6 empty] | TT_K(2)=SQR(SQR(TT_K(2)))
'Convert target temperature into desired units
TT_C( | 2 | )=TT_K( | 2 | )-273.15 | [6 empty] | TT_C(2)=TT_K(2)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 3 | 3 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 3 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(3),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 3 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(3),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(3)= | 1392520000 | +( | 7230830 | *SBT_C(3))+( | 75298.8 | *SBT_C(3)^2) |  |  |  |  | m(3)=1392520000+(7230830*SBT_C(3))+(75298.8*SBT_C(3)^2)
b(3)= | -2209980 | +( | -62020 | *SBT_C(3))+( | 14222.2 | *SBT_C(3)^2) |  |  |  |  | b(3)=-2209980+(-62020*SBT_C(3))+(14222.2*SBT_C(3)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 3 | )=SBT_C( | 3 | )+273.15 | [6 empty] | SBT_K(3)=SBT_C(3)+273.15
TT_K( | 3 | )=SBT_K( | 3 | )^4+TTmV(3) | *m(3) | +b(3) |  |  |  |  | TT_K(3)=SBT_K(3)^4+TTmV(3)*m(3)+b(3)
TT_K( | 3 | )=SQR(SQR(TT_K( | 3 | ))) | [6 empty] | TT_K(3)=SQR(SQR(TT_K(3)))
'Convert target temperature into desired units
TT_C( | 3 | )=TT_K( | 3 | )-273.15 | [6 empty] | TT_C(3)=TT_K(3)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 4 | 4 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 4 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(4),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 4 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(4),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(4)= | 1341860000 | +( | 7045270 | *SBT_C(4))+( | 73401.4 | *SBT_C(4)^2) |  |  |  |  | m(4)=1341860000+(7045270*SBT_C(4))+(73401.4*SBT_C(4)^2)
b(4)= | 647879 | +( | 24049.6 | *SBT_C(4))+( | 14900.8 | *SBT_C(4)^2) |  |  |  |  | b(4)=647879+(24049.6*SBT_C(4))+(14900.8*SBT_C(4)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 4 | )=SBT_C( | 4 | )+273.15 | [6 empty] | SBT_K(4)=SBT_C(4)+273.15
TT_K( | 4 | )=SBT_K( | 4 | )^4+TTmV(4) | *m(4) | +b(4) |  |  |  |  | TT_K(4)=SBT_K(4)^4+TTmV(4)*m(4)+b(4)
TT_K( | 4 | )=SQR(SQR(TT_K( | 4 | ))) | [6 empty] | TT_K(4)=SQR(SQR(TT_K(4)))
'Convert target temperature into desired units
TT_C( | 4 | )=TT_K( | 4 | )-273.15 | [6 empty] | TT_C(4)=TT_K(4)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 5 | 5 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 5 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(5),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 5 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(5),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(5)= | 1386130000 | +( | 7242140 | *SBT_C(5))+( | 70955 | *SBT_C(5)^2) |  |  |  |  | m(5)=1386130000+(7242140*SBT_C(5))+(70955*SBT_C(5)^2)
b(5)= | -2918440 | +( | -11589.2 | *SBT_C(5))+( | 14149.3 | *SBT_C(5)^2) |  |  |  |  | b(5)=-2918440+(-11589.2*SBT_C(5))+(14149.3*SBT_C(5)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 5 | )=SBT_C( | 5 | )+273.15 | [6 empty] | SBT_K(5)=SBT_C(5)+273.15
TT_K( | 5 | )=SBT_K( | 5 | )^4+TTmV(5) | *m(5) | +b(5) |  |  |  |  | TT_K(5)=SBT_K(5)^4+TTmV(5)*m(5)+b(5)
TT_K( | 5 | )=SQR(SQR(TT_K( | 5 | ))) | [6 empty] | TT_K(5)=SQR(SQR(TT_K(5)))
'Convert target temperature into desired units
TT_C( | 5 | )=TT_K( | 5 | )-273.15 | [6 empty] | TT_C(5)=TT_K(5)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 6 | 6 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 6 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(6),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 6 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(6),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(6)= | 1368500000 | +( | 7282460 | *SBT_C(6))+( | 75238.7 | *SBT_C(6)^2) |  |  |  |  | m(6)=1368500000+(7282460*SBT_C(6))+(75238.7*SBT_C(6)^2)
b(6)= | -5148230 | +( | 94378.8 | *SBT_C(6))+( | 14931.7 | *SBT_C(6)^2) |  |  |  |  | b(6)=-5148230+(94378.8*SBT_C(6))+(14931.7*SBT_C(6)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 6 | )=SBT_C( | 6 | )+273.15 | [6 empty] | SBT_K(6)=SBT_C(6)+273.15
TT_K( | 6 | )=SBT_K( | 6 | )^4+TTmV(6) | *m(6) | +b(6) |  |  |  |  | TT_K(6)=SBT_K(6)^4+TTmV(6)*m(6)+b(6)
TT_K( | 6 | )=SQR(SQR(TT_K( | 6 | ))) | [6 empty] | TT_K(6)=SQR(SQR(TT_K(6)))
'Convert target temperature into desired units
TT_C( | 6 | )=TT_K( | 6 | )-273.15 | [6 empty] | TT_C(6)=TT_K(6)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 7 | 7 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 7 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(7),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 7 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(7),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(7)= | 1371730000 | +( | 7058920 | *SBT_C(7))+( | 76906 | *SBT_C(7)^2) |  |  |  |  | m(7)=1371730000+(7058920*SBT_C(7))+(76906*SBT_C(7)^2)
b(7)= | -6107980 | +( | -35518.4 | *SBT_C(7))+( | 14153.7 | *SBT_C(7)^2) |  |  |  |  | b(7)=-6107980+(-35518.4*SBT_C(7))+(14153.7*SBT_C(7)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 7 | )=SBT_C( | 7 | )+273.15 | [6 empty] | SBT_K(7)=SBT_C(7)+273.15
TT_K( | 7 | )=SBT_K( | 7 | )^4+TTmV(7) | *m(7) | +b(7) |  |  |  |  | TT_K(7)=SBT_K(7)^4+TTmV(7)*m(7)+b(7)
TT_K( | 7 | )=SQR(SQR(TT_K( | 7 | ))) | [6 empty] | TT_K(7)=SQR(SQR(TT_K(7)))
'Convert target temperature into desired units
TT_C( | 7 | )=TT_K( | 7 | )-273.15 | [6 empty] | TT_C(7)=TT_K(7)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 8 | 8 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 8 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(8),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 8 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(8),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(8)= | 1434740000 | +( | 7449330 | *SBT_C(8))+( | 73361.9 | *SBT_C(8)^2) |  |  |  |  | m(8)=1434740000+(7449330*SBT_C(8))+(73361.9*SBT_C(8)^2)
b(8)= | -3427590 | +( | 28290.2 | *SBT_C(8))+( | 15814.1 | *SBT_C(8)^2) |  |  |  |  | b(8)=-3427590+(28290.2*SBT_C(8))+(15814.1*SBT_C(8)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 8 | )=SBT_C( | 8 | )+273.15 | [6 empty] | SBT_K(8)=SBT_C(8)+273.15
TT_K( | 8 | )=SBT_K( | 8 | )^4+TTmV(8) | *m(8) | +b(8) |  |  |  |  | TT_K(8)=SBT_K(8)^4+TTmV(8)*m(8)+b(8)
TT_K( | 8 | )=SQR(SQR(TT_K( | 8 | ))) | [6 empty] | TT_K(8)=SQR(SQR(TT_K(8)))
'Convert target temperature into desired units
TT_C( | 8 | )=TT_K( | 8 | )-273.15 | [6 empty] | TT_C(8)=TT_K(8)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 9 | 9 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 9 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(9),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 9 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(9),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(9)= | 1367620000 | +( | 7237750 | *SBT_C(9))+( | 71645.9 | *SBT_C(9)^2) |  |  |  |  | m(9)=1367620000+(7237750*SBT_C(9))+(71645.9*SBT_C(9)^2)
b(9)= | -5766930 | +( | -21127.4 | *SBT_C(9))+( | 14807.4 | *SBT_C(9)^2) |  |  |  |  | b(9)=-5766930+(-21127.4*SBT_C(9))+(14807.4*SBT_C(9)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 9 | )=SBT_C( | 9 | )+273.15 | [6 empty] | SBT_K(9)=SBT_C(9)+273.15
TT_K( | 9 | )=SBT_K( | 9 | )^4+TTmV(9) | *m(9) | +b(9) |  |  |  |  | TT_K(9)=SBT_K(9)^4+TTmV(9)*m(9)+b(9)
TT_K( | 9 | )=SQR(SQR(TT_K( | 9 | ))) | [6 empty] | TT_K(9)=SQR(SQR(TT_K(9)))
'Convert target temperature into desired units
TT_C( | 9 | )=TT_K( | 9 | )-273.15 | [6 empty] | TT_C(9)=TT_K(9)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 10 | 10 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 10 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(10),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 10 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(10),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(10)= | 1339840000 | +( | 7046660 | *SBT_C(10))+( | 74491.8 | *SBT_C(10)^2) |  |  |  |  | m(10)=1339840000+(7046660*SBT_C(10))+(74491.8*SBT_C(10)^2)
b(10)= | -5004960 | +( | 57560.6 | *SBT_C(10))+( | 13317.6 | *SBT_C(10)^2) |  |  |  |  | b(10)=-5004960+(57560.6*SBT_C(10))+(13317.6*SBT_C(10)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 10 | )=SBT_C( | 10 | )+273.15 | [6 empty] | SBT_K(10)=SBT_C(10)+273.15
TT_K( | 10 | )=SBT_K( | 10 | )^4+TTmV(10) | *m(10) | +b(10) |  |  |  |  | TT_K(10)=SBT_K(10)^4+TTmV(10)*m(10)+b(10)
TT_K( | 10 | )=SQR(SQR(TT_K( | 10 | ))) | [6 empty] | TT_K(10)=SQR(SQR(TT_K(10)))
'Convert target temperature into desired units
TT_C( | 10 | )=TT_K( | 10 | )-273.15 | [6 empty] | TT_C(10)=TT_K(10)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 11 | 11 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 11 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(11),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 11 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(11),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(11)= | 1737340000 | +( | 9126780 | *SBT_C(11))+( | 95247.2 | *SBT_C(11)^2) |  |  |  |  | m(11)=1737340000+(9126780*SBT_C(11))+(95247.2*SBT_C(11)^2)
b(11)= | -18151800 | +( | 888.145 | *SBT_C(11))+( | 15603.9 | *SBT_C(11)^2) |  |  |  |  | b(11)=-18151800+(888.145*SBT_C(11))+(15603.9*SBT_C(11)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 11 | )=SBT_C( | 11 | )+273.15 | [6 empty] | SBT_K(11)=SBT_C(11)+273.15
TT_K( | 11 | )=SBT_K( | 11 | )^4+TTmV(11) | *m(11) | +b(11) |  |  |  |  | TT_K(11)=SBT_K(11)^4+TTmV(11)*m(11)+b(11)
TT_K( | 11 | )=SQR(SQR(TT_K( | 11 | ))) | [6 empty] | TT_K(11)=SQR(SQR(TT_K(11)))
'Convert target temperature into desired units
TT_C( | 11 | )=TT_K( | 11 | )-273.15 | [6 empty] | TT_C(11)=TT_K(11)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 12 | 12 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 12 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(12),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 12 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(12),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(12)= | 1331080000 | +( | 6996190 | *SBT_C(12))+( | 72127.4 | *SBT_C(12)^2) |  |  |  |  | m(12)=1331080000+(6996190*SBT_C(12))+(72127.4*SBT_C(12)^2)
b(12)= | -8744260 | +( | -40772.5 | *SBT_C(12))+( | 14361 | *SBT_C(12)^2) |  |  |  |  | b(12)=-8744260+(-40772.5*SBT_C(12))+(14361*SBT_C(12)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 12 | )=SBT_C( | 12 | )+273.15 | [6 empty] | SBT_K(12)=SBT_C(12)+273.15
TT_K( | 12 | )=SBT_K( | 12 | )^4+TTmV(12) | *m(12) | +b(12) |  |  |  |  | TT_K(12)=SBT_K(12)^4+TTmV(12)*m(12)+b(12)
TT_K( | 12 | )=SQR(SQR(TT_K( | 12 | ))) | [6 empty] | TT_K(12)=SQR(SQR(TT_K(12)))
'Convert target temperature into desired units
TT_C( | 12 | )=TT_K( | 12 | )-273.15 | [6 empty] | TT_C(12)=TT_K(12)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 13 | 13 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 13 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(13),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 13 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(13),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(13)= | 1395890000 | +( | 7277220 | *SBT_C(13))+( | 75992 | *SBT_C(13)^2) |  |  |  |  | m(13)=1395890000+(7277220*SBT_C(13))+(75992*SBT_C(13)^2)
b(13)= | -2662950 | +( | -9950.63 | *SBT_C(13))+( | 13555.7 | *SBT_C(13)^2) |  |  |  |  | b(13)=-2662950+(-9950.63*SBT_C(13))+(13555.7*SBT_C(13)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 13 | )=SBT_C( | 13 | )+273.15 | [6 empty] | SBT_K(13)=SBT_C(13)+273.15
TT_K( | 13 | )=SBT_K( | 13 | )^4+TTmV(13) | *m(13) | +b(13) |  |  |  |  | TT_K(13)=SBT_K(13)^4+TTmV(13)*m(13)+b(13)
TT_K( | 13 | )=SQR(SQR(TT_K( | 13 | ))) | [6 empty] | TT_K(13)=SQR(SQR(TT_K(13)))
'Convert target temperature into desired units
TT_C( | 13 | )=TT_K( | 13 | )-273.15 | [6 empty] | TT_C(13)=TT_K(13)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 14 | 14 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 14 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(14),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 14 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(14),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(14)= | 1373910000 | +( | 7204930 | *SBT_C(14))+( | 71235.4 | *SBT_C(14)^2) |  |  |  |  | m(14)=1373910000+(7204930*SBT_C(14))+(71235.4*SBT_C(14)^2)
b(14)= | 2801350 | +( | -42316.6 | *SBT_C(14))+( | 14731.4 | *SBT_C(14)^2) |  |  |  |  | b(14)=2801350+(-42316.6*SBT_C(14))+(14731.4*SBT_C(14)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 14 | )=SBT_C( | 14 | )+273.15 | [6 empty] | SBT_K(14)=SBT_C(14)+273.15
TT_K( | 14 | )=SBT_K( | 14 | )^4+TTmV(14) | *m(14) | +b(14) |  |  |  |  | TT_K(14)=SBT_K(14)^4+TTmV(14)*m(14)+b(14)
TT_K( | 14 | )=SQR(SQR(TT_K( | 14 | ))) | [6 empty] | TT_K(14)=SQR(SQR(TT_K(14)))
'Convert target temperature into desired units
TT_C( | 14 | )=TT_K( | 14 | )-273.15 | [6 empty] | TT_C(14)=TT_K(14)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 15 | 15 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 15 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(15),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 15 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(15),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(15)= | 1410870000 | +( | 7449540 | *SBT_C(15))+( | 77553.2 | *SBT_C(15)^2) |  |  |  |  | m(15)=1410870000+(7449540*SBT_C(15))+(77553.2*SBT_C(15)^2)
b(15)= | -4211160 | +( | 16325.6 | *SBT_C(15))+( | 12858.8 | *SBT_C(15)^2) |  |  |  |  | b(15)=-4211160+(16325.6*SBT_C(15))+(12858.8*SBT_C(15)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 15 | )=SBT_C( | 15 | )+273.15 | [6 empty] | SBT_K(15)=SBT_C(15)+273.15
TT_K( | 15 | )=SBT_K( | 15 | )^4+TTmV(15) | *m(15) | +b(15) |  |  |  |  | TT_K(15)=SBT_K(15)^4+TTmV(15)*m(15)+b(15)
TT_K( | 15 | )=SQR(SQR(TT_K( | 15 | ))) | [6 empty] | TT_K(15)=SQR(SQR(TT_K(15)))
'Convert target temperature into desired units
TT_C( | 15 | )=TT_K( | 15 | )-273.15 | [6 empty] | TT_C(15)=TT_K(15)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
1 | 16 | 16 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 16 | ),1, | 3 | , | VX2 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(16),1,3,VX2,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 16 | ),1,mV2_5, | 1 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(16),1,mV2_5,1,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(16)= | 1321880000 | +( | 6953780 | *SBT_C(16))+( | 75548.9 | *SBT_C(16)^2) |  |  |  |  | m(16)=1321880000+(6953780*SBT_C(16))+(75548.9*SBT_C(16)^2)
b(16)= | -7079850 | +( | 128729 | *SBT_C(16))+( | 13473.3 | *SBT_C(16)^2) |  |  |  |  | b(16)=-7079850+(128729*SBT_C(16))+(13473.3*SBT_C(16)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 16 | )=SBT_C( | 16 | )+273.15 | [6 empty] | SBT_K(16)=SBT_C(16)+273.15
TT_K( | 16 | )=SBT_K( | 16 | )^4+TTmV(16) | *m(16) | +b(16) |  |  |  |  | TT_K(16)=SBT_K(16)^4+TTmV(16)*m(16)+b(16)
TT_K( | 16 | )=SQR(SQR(TT_K( | 16 | ))) | [6 empty] | TT_K(16)=SQR(SQR(TT_K(16)))
'Convert target temperature into desired units
TT_C( | 16 | )=TT_K( | 16 | )-273.15 | [6 empty] | TT_C(16)=TT_K(16)-273.15
'Turn AM16/32 Multiplexer Off. PortSet(DataLogger channel C#, State 0=off and reset)
PortSet( | 2 | ,0) | [6 empty] | PortSet(2,0)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MUXNumber | 2 | IR sensors, thermocouples and heat flux | [6 empty] | 'Turn AM16/32 Multiplexer On. PortSet(DataLogger com channel, State 1=on and can respond to clock pulses)
PortSet( | 3 | ,1) | [6 empty] | PortSet(3,1)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 1 | 17 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 17 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(17),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 17 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(17),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(17)= | 1347710000 | +( | 7097850 | *SBT_C(17))+( | 73741.2 | *SBT_C(17)^2) |  |  |  |  | m(17)=1347710000+(7097850*SBT_C(17))+(73741.2*SBT_C(17)^2)
b(17)= | -5754910 | +( | -18618.3 | *SBT_C(17))+( | 13883.8 | *SBT_C(17)^2) |  |  |  |  | b(17)=-5754910+(-18618.3*SBT_C(17))+(13883.8*SBT_C(17)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 17 | )=SBT_C( | 17 | )+273.15 | [6 empty] | SBT_K(17)=SBT_C(17)+273.15
TT_K( | 17 | )=SBT_K( | 17 | )^4+TTmV(17) | *m(17) | +b(17) |  |  |  |  | TT_K(17)=SBT_K(17)^4+TTmV(17)*m(17)+b(17)
TT_K( | 17 | )=SQR(SQR(TT_K( | 17 | ))) | [6 empty] | TT_K(17)=SQR(SQR(TT_K(17)))
'Convert target temperature into desired units
TT_C( | 17 | )=TT_K( | 17 | )-273.15 | [6 empty] | TT_C(17)=TT_K(17)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 2 | 18 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 18 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(18),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 18 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(18),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(18)= | 1340100000 | +( | 6959910 | *SBT_C(18))+( | 74279.9 | *SBT_C(18)^2) |  |  |  |  | m(18)=1340100000+(6959910*SBT_C(18))+(74279.9*SBT_C(18)^2)
b(18)= | -683085 | +( | 70789.2 | *SBT_C(18))+( | 13195.5 | *SBT_C(18)^2) |  |  |  |  | b(18)=-683085+(70789.2*SBT_C(18))+(13195.5*SBT_C(18)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 18 | )=SBT_C( | 18 | )+273.15 | [6 empty] | SBT_K(18)=SBT_C(18)+273.15
TT_K( | 18 | )=SBT_K( | 18 | )^4+TTmV(18) | *m(18) | +b(18) |  |  |  |  | TT_K(18)=SBT_K(18)^4+TTmV(18)*m(18)+b(18)
TT_K( | 18 | )=SQR(SQR(TT_K( | 18 | ))) | [6 empty] | TT_K(18)=SQR(SQR(TT_K(18)))
'Convert target temperature into desired units
TT_C( | 18 | )=TT_K( | 18 | )-273.15 | [6 empty] | TT_C(18)=TT_K(18)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 3 | 19 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 19 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(19),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 19 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(19),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(19)= | 1321320000 | +( | 6942310 | *SBT_C(19))+( | 77904.1 | *SBT_C(19)^2) |  |  |  |  | m(19)=1321320000+(6942310*SBT_C(19))+(77904.1*SBT_C(19)^2)
b(19)= | -9474180 | +( | 85649.1 | *SBT_C(19))+( | 12783 | *SBT_C(19)^2) |  |  |  |  | b(19)=-9474180+(85649.1*SBT_C(19))+(12783*SBT_C(19)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 19 | )=SBT_C( | 19 | )+273.15 | [6 empty] | SBT_K(19)=SBT_C(19)+273.15
TT_K( | 19 | )=SBT_K( | 19 | )^4+TTmV(19) | *m(19) | +b(19) |  |  |  |  | TT_K(19)=SBT_K(19)^4+TTmV(19)*m(19)+b(19)
TT_K( | 19 | )=SQR(SQR(TT_K( | 19 | ))) | [6 empty] | TT_K(19)=SQR(SQR(TT_K(19)))
'Convert target temperature into desired units
TT_C( | 19 | )=TT_K( | 19 | )-273.15 | [6 empty] | TT_C(19)=TT_K(19)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 4 | 20 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 20 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(20),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 20 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(20),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(20)= | 1414580000 | +( | 7496330 | *SBT_C(20))+( | 79768.7 | *SBT_C(20)^2) |  |  |  |  | m(20)=1414580000+(7496330*SBT_C(20))+(79768.7*SBT_C(20)^2)
b(20)= | -4488640 | +( | 127752 | *SBT_C(20))+( | 14089.7 | *SBT_C(20)^2) |  |  |  |  | b(20)=-4488640+(127752*SBT_C(20))+(14089.7*SBT_C(20)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 20 | )=SBT_C( | 20 | )+273.15 | [6 empty] | SBT_K(20)=SBT_C(20)+273.15
TT_K( | 20 | )=SBT_K( | 20 | )^4+TTmV(20) | *m(20) | +b(20) |  |  |  |  | TT_K(20)=SBT_K(20)^4+TTmV(20)*m(20)+b(20)
TT_K( | 20 | )=SQR(SQR(TT_K( | 20 | ))) | [6 empty] | TT_K(20)=SQR(SQR(TT_K(20)))
'Convert target temperature into desired units
TT_C( | 20 | )=TT_K( | 20 | )-273.15 | [6 empty] | TT_C(20)=TT_K(20)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 5 | 21 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 21 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(21),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 21 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(21),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(21)= | 1367620000 | +( | 7237750 | *SBT_C(21))+( | 71645.9 | *SBT_C(21)^2) |  |  |  |  | m(21)=1367620000+(7237750*SBT_C(21))+(71645.9*SBT_C(21)^2)
b(21)= | -5766930 | +( | -21127.4 | *SBT_C(21))+( | 14807.4 | *SBT_C(21)^2) |  |  |  |  | b(21)=-5766930+(-21127.4*SBT_C(21))+(14807.4*SBT_C(21)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 21 | )=SBT_C( | 21 | )+273.15 | [6 empty] | SBT_K(21)=SBT_C(21)+273.15
TT_K( | 21 | )=SBT_K( | 21 | )^4+TTmV(21) | *m(21) | +b(21) |  |  |  |  | TT_K(21)=SBT_K(21)^4+TTmV(21)*m(21)+b(21)
TT_K( | 21 | )=SQR(SQR(TT_K( | 21 | ))) | [6 empty] | TT_K(21)=SQR(SQR(TT_K(21)))
'Convert target temperature into desired units
TT_C( | 21 | )=TT_K( | 21 | )-273.15 | [6 empty] | TT_C(21)=TT_K(21)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 6 | 22 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 22 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(22),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 22 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(22),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(22)= | 1332630000 | +( | 7031600 | *SBT_C(22))+( | 73620.1 | *SBT_C(22)^2) |  |  |  |  | m(22)=1332630000+(7031600*SBT_C(22))+(73620.1*SBT_C(22)^2)
b(22)= | -1572190 | +( | 37528.5 | *SBT_C(22))+( | 12417.8 | *SBT_C(22)^2) |  |  |  |  | b(22)=-1572190+(37528.5*SBT_C(22))+(12417.8*SBT_C(22)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 22 | )=SBT_C( | 22 | )+273.15 | [6 empty] | SBT_K(22)=SBT_C(22)+273.15
TT_K( | 22 | )=SBT_K( | 22 | )^4+TTmV(22) | *m(22) | +b(22) |  |  |  |  | TT_K(22)=SBT_K(22)^4+TTmV(22)*m(22)+b(22)
TT_K( | 22 | )=SQR(SQR(TT_K( | 22 | ))) | [6 empty] | TT_K(22)=SQR(SQR(TT_K(22)))
'Convert target temperature into desired units
TT_C( | 22 | )=TT_K( | 22 | )-273.15 | [6 empty] | TT_C(22)=TT_K(22)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 7 | 23 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 23 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(23),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 23 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(23),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(23)= | 1432990000 | +( | 7521750 | *SBT_C(23))+( | 78210.1 | *SBT_C(23)^2) |  |  |  |  | m(23)=1432990000+(7521750*SBT_C(23))+(78210.1*SBT_C(23)^2)
b(23)= | -3974070 | +( | 3509.1 | *SBT_C(23))+( | 13660.6 | *SBT_C(23)^2) |  |  |  |  | b(23)=-3974070+(3509.1*SBT_C(23))+(13660.6*SBT_C(23)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 23 | )=SBT_C( | 23 | )+273.15 | [6 empty] | SBT_K(23)=SBT_C(23)+273.15
TT_K( | 23 | )=SBT_K( | 23 | )^4+TTmV(23) | *m(23) | +b(23) |  |  |  |  | TT_K(23)=SBT_K(23)^4+TTmV(23)*m(23)+b(23)
TT_K( | 23 | )=SQR(SQR(TT_K( | 23 | ))) | [6 empty] | TT_K(23)=SQR(SQR(TT_K(23)))
'Convert target temperature into desired units
TT_C( | 23 | )=TT_K( | 23 | )-273.15 | [6 empty] | TT_C(23)=TT_K(23)-273.15
MUXNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 | 8 | 24 | IR sensor | [6 empty] | PulsePort(1,10000)
'SI-111 Precision Infrared Radiometer measurements 'TT_C' and 'SBT_C'
'Measure SI-111 sensor body thermistor temperature. Therm109 (Dest, Reps, SEChan, VX/ExChan, SettlingTime, Integ, Mult, Offset)
Therm109(SBT_C( | 24 | ),1, | 4 | , | VX3 | ,0,_60Hz,1,0) |  |  |  |  | Therm109(SBT_C(24),1,4,VX3,0,_60Hz,1,0)
'Measure SI-111 output of thermopile. VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(TTmV( | 24 | ),1,mV2_5, | 3 | ,True,0,_60Hz,1,0,) | [6 empty] | VoltDiff(TTmV(24),1,mV2_5,3,True,0,_60Hz,1,0,)
| 6 |  | 7 |  | 8 | [5 empty] | 'Calculate slope (m) and offset (b) coefficients for target temperature calculation
m(24)= | 1432430000 | +( | 7402560 | *SBT_C(24))+( | 79259.3 | *SBT_C(24)^2) |  |  |  |  | m(24)=1432430000+(7402560*SBT_C(24))+(79259.3*SBT_C(24)^2)
b(24)= | -6857680 | +( | 133432 | *SBT_C(24))+( | 13799.7 | *SBT_C(24)^2) |  |  |  |  | b(24)=-6857680+(133432*SBT_C(24))+(13799.7*SBT_C(24)^2)
'Calculate target temperature using calculated slope (m) and offset (b)
SBT_K( | 24 | )=SBT_C( | 24 | )+273.15 | [6 empty] | SBT_K(24)=SBT_C(24)+273.15
TT_K( | 24 | )=SBT_K( | 24 | )^4+TTmV(24) | *m(24) | +b(24) |  |  |  |  | TT_K(24)=SBT_K(24)^4+TTmV(24)*m(24)+b(24)
TT_K( | 24 | )=SQR(SQR(TT_K( | 24 | ))) | [6 empty] | TT_K(24)=SQR(SQR(TT_K(24)))
'Convert target temperature into desired units
TT_C( | 24 | )=TT_K( | 24 | )-273.15 | [6 empty] | TT_C(24)=TT_K(24)-273.15
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 1 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 1 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(1),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 1 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.39 | ,0) |  |  |  |  | VoltSe (HeatFlux(1),1,MV7_5,4,0,0,_60Hz,16.39,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 2 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 2 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(2),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 2 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.1 | ,0) |  |  |  |  | VoltSe (HeatFlux(2),1,MV7_5,4,0,0,_60Hz,16.1,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 3 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 3 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(3),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 3 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.71 | ,0) |  |  |  |  | VoltSe (HeatFlux(3),1,MV7_5,4,0,0,_60Hz,16.71,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 4 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 4 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(4),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 4 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.53 | ,0) |  |  |  |  | VoltSe (HeatFlux(4),1,MV7_5,4,0,0,_60Hz,16.53,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 5 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 5 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(5),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 5 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.39 | ,0) |  |  |  |  | VoltSe (HeatFlux(5),1,MV7_5,4,0,0,_60Hz,16.39,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 6 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 6 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(6),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 6 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.08 | ,0) |  |  |  |  | VoltSe (HeatFlux(6),1,MV7_5,4,0,0,_60Hz,16.08,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 7 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 7 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(7),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 7 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.14 | ,0) |  |  |  |  | VoltSe (HeatFlux(7),1,MV7_5,4,0,0,_60Hz,16.14,0)
MuxNumber | MUX Chanel (4X14 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
2 |  | 8 | Thermocouple & Heat Flux | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 8 | ),1,MV2_5C, | 3 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(8),1,MV2_5C,3,TypeE,PTemp_C,True,0,_60Hz,1,0)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 8 | ),1,MV7_5, | 4 | ,0,0,_60Hz, | 16.14 | ,0) |  |  |  |  | VoltSe (HeatFlux(8),1,MV7_5,4,0,0,_60Hz,16.14,0)
'Turn AM16/32 Multiplexer Off. PortSet(DataLogger channel C#, State 0=off and reset)
PortSet( | 3 | ,0) | [6 empty] | PortSet(3,0)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MuxNumber | 3 | Radiation reflection and soil temp | [6 empty] | 'Turn AM16/32 Multiplexer On. PortSet(DataLogger com channel, State 1=on and can respond to clock pulses)
PortSet( | 4 | ,1) | [6 empty] | PortSet(4,1)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 1 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 1 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(1),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 1 | ) < 0 Then SlrReflectkW( | 1 | ) = 0 | [6 empty] | if SlrReflectkW(1) < 0 Then SlrReflectkW(1) = 0
SlrReflectMJ( | 1 | )=SlrReflectkW( | 1 | ) * | 0.12 | [5 empty] | SlrReflectMJ(1)=SlrReflectkW(1) * 0.12
SlrReflectkW( | 1 | )=SlrReflectkW( | 1 | ) * | 0.2 | [5 empty] | SlrReflectkW(1)=SlrReflectkW(1) * 0.2
|  | 24 | Soil surface temperature | [7 empty] | '108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 24 | ),1, | 16 | , | VX1 | ,0,_60Hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(24),1,16,VX1,0,_60Hz,1,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 2 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 2 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(2),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 2 | ) < 0 Then SlrReflectkW( | 2 | ) = 0 | [6 empty] | if SlrReflectkW(2) < 0 Then SlrReflectkW(2) = 0
SlrReflectMJ( | 2 | )=SlrReflectkW( | 2 | ) * | 0.12 | [5 empty] | SlrReflectMJ(2)=SlrReflectkW(2) * 0.12
SlrReflectkW( | 2 | )=SlrReflectkW( | 2 | ) * | 0.2 | [5 empty] | SlrReflectkW(2)=SlrReflectkW(2) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 3 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 3 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(3),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 3 | ) < 0 Then SlrReflectkW( | 3 | ) = 0 | [6 empty] | if SlrReflectkW(3) < 0 Then SlrReflectkW(3) = 0
SlrReflectMJ( | 3 | )=SlrReflectkW( | 3 | ) * | 0.12 | [5 empty] | SlrReflectMJ(3)=SlrReflectkW(3) * 0.12
SlrReflectkW( | 3 | )=SlrReflectkW( | 3 | ) * | 0.2 | [5 empty] | SlrReflectkW(3)=SlrReflectkW(3) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 4 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 4 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(4),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 4 | ) < 0 Then SlrReflectkW( | 4 | ) = 0 | [6 empty] | if SlrReflectkW(4) < 0 Then SlrReflectkW(4) = 0
SlrReflectMJ( | 4 | )=SlrReflectkW( | 4 | ) * | 0.12 | [5 empty] | SlrReflectMJ(4)=SlrReflectkW(4) * 0.12
SlrReflectkW( | 4 | )=SlrReflectkW( | 4 | ) * | 0.2 | [5 empty] | SlrReflectkW(4)=SlrReflectkW(4) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 5 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 5 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(5),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 5 | ) < 0 Then SlrReflectkW( | 5 | ) = 0 | [6 empty] | if SlrReflectkW(5) < 0 Then SlrReflectkW(5) = 0
SlrReflectMJ( | 5 | )=SlrReflectkW( | 5 | ) * | 0.12 | [5 empty] | SlrReflectMJ(5)=SlrReflectkW(5) * 0.12
SlrReflectkW( | 5 | )=SlrReflectkW( | 5 | ) * | 0.2 | [5 empty] | SlrReflectkW(5)=SlrReflectkW(5) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 6 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 6 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(6),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 6 | ) < 0 Then SlrReflectkW( | 6 | ) = 0 | [6 empty] | if SlrReflectkW(6) < 0 Then SlrReflectkW(6) = 0
SlrReflectMJ( | 6 | )=SlrReflectkW( | 6 | ) * | 0.12 | [5 empty] | SlrReflectMJ(6)=SlrReflectkW(6) * 0.12
SlrReflectkW( | 6 | )=SlrReflectkW( | 6 | ) * | 0.2 | [5 empty] | SlrReflectkW(6)=SlrReflectkW(6) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 7 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 7 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(7),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 7 | ) < 0 Then SlrReflectkW( | 7 | ) = 0 | [6 empty] | if SlrReflectkW(7) < 0 Then SlrReflectkW(7) = 0
SlrReflectMJ( | 7 | )=SlrReflectkW( | 7 | ) * | 0.12 | [5 empty] | SlrReflectMJ(7)=SlrReflectkW(7) * 0.12
SlrReflectkW( | 7 | )=SlrReflectkW( | 7 | ) * | 0.2 | [5 empty] | SlrReflectkW(7)=SlrReflectkW(7) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 8 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 8 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(8),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 8 | ) < 0 Then SlrReflectkW( | 8 | ) = 0 | [6 empty] | if SlrReflectkW(8) < 0 Then SlrReflectkW(8) = 0
SlrReflectMJ( | 8 | )=SlrReflectkW( | 8 | ) * | 0.12 | [5 empty] | SlrReflectMJ(8)=SlrReflectkW(8) * 0.12
SlrReflectkW( | 8 | )=SlrReflectkW( | 8 | ) * | 0.2 | [5 empty] | SlrReflectkW(8)=SlrReflectkW(8) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 9 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 9 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(9),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 9 | ) < 0 Then SlrReflectkW( | 9 | ) = 0 | [6 empty] | if SlrReflectkW(9) < 0 Then SlrReflectkW(9) = 0
SlrReflectMJ( | 9 | )=SlrReflectkW( | 9 | ) * | 0.12 | [5 empty] | SlrReflectMJ(9)=SlrReflectkW(9) * 0.12
SlrReflectkW( | 9 | )=SlrReflectkW( | 9 | ) * | 0.2 | [5 empty] | SlrReflectkW(9)=SlrReflectkW(9) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 10 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 10 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(10),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 10 | ) < 0 Then SlrReflectkW( | 10 | ) = 0 | [6 empty] | if SlrReflectkW(10) < 0 Then SlrReflectkW(10) = 0
SlrReflectMJ( | 10 | )=SlrReflectkW( | 10 | ) * | 0.12 | [5 empty] | SlrReflectMJ(10)=SlrReflectkW(10) * 0.12
SlrReflectkW( | 10 | )=SlrReflectkW( | 10 | ) * | 0.2 | [5 empty] | SlrReflectkW(10)=SlrReflectkW(10) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 11 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 11 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(11),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 11 | ) < 0 Then SlrReflectkW( | 11 | ) = 0 | [6 empty] | if SlrReflectkW(11) < 0 Then SlrReflectkW(11) = 0
SlrReflectMJ( | 11 | )=SlrReflectkW( | 11 | ) * | 0.12 | [5 empty] | SlrReflectMJ(11)=SlrReflectkW(11) * 0.12
SlrReflectkW( | 11 | )=SlrReflectkW( | 11 | ) * | 0.2 | [5 empty] | SlrReflectkW(11)=SlrReflectkW(11) * 0.2
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
3 |  | 12 | Reflected Radn | [6 empty] | PulsePort(1,10000)
'LI200X Pyranometer measurements 'SlrMJReflect' and 'SlrkWReflect' VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(SlrReflectkW( | 12 | ),1,mV7_5, | 4 | ,True,0,_60Hz,1,0) | [6 empty] | VoltDiff(SlrReflectkW(12),1,mV7_5,4,True,0,_60Hz,1,0)
if SlrReflectkW( | 12 | ) < 0 Then SlrReflectkW( | 12 | ) = 0 | [6 empty] | if SlrReflectkW(12) < 0 Then SlrReflectkW(12) = 0
SlrReflectMJ( | 12 | )=SlrReflectkW( | 12 | ) * | 0.12 | [5 empty] | SlrReflectMJ(12)=SlrReflectkW(12) * 0.12
SlrReflectkW( | 12 | )=SlrReflectkW( | 12 | ) * | 0.2 | [5 empty] | SlrReflectkW(12)=SlrReflectkW(12) * 0.2
'Turn AM16/32 Multiplexer Off. PortSet(DataLogger channel C#, State 0=off and reset)
PortSet( | 4 | ,0) | [6 empty] | PortSet(4,0)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MuxNumber | 4 | Soil Temperature and Heat Flux | [6 empty] | 'Turn AM16/32 Multiplexer On. PortSet(DataLogger com channel, State 1=on and can respond to clock pulses)
PortSet( | 5 | ,1) | [6 empty] | PortSet(5,1)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 1 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 1 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(1),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 2 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 2 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(2),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 3 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 3 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(3),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 4 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 4 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(4),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 5 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 5 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(5),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 6 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 6 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(6),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 7 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 7 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(7),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 8 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 8 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(8),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 9 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 9 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(9),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 10 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 10 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(10),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 11 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 11 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(11),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 12 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 12 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(12),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 13 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 13 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(13),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 14 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 14 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(14),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 15 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 15 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(15),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 16 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 16 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(16),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 17 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 17 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(17),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 18 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 18 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(18),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 19 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 19 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(19),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 20 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 20 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(20),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 21 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 21 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(21),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 22 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 22 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(22),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 23 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'108 Temperature Probe measurement 'T108_C. Therm108(Dest,Reps,SEChan,VxChan,Settlingtime,Integration,Multiplier,Offset)
Therm108(SoilSurfTC( | 23 | ),1, | 16 | , | VX1 | ,0,_60hz,1,0) |  |  |  |  | Therm108(SoilSurfTC(23),1,16,VX1,0,_60hz,1,0)
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 24 | Soil surface temperature | [6 empty] | PulsePort(1,10000)
'port not working, skip
MUXNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
4 |  | 9 | Heat Flux | [6 empty] | PulsePort(1,10000)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 9 | ),1,MV7_5, | 16 | ,0,0,_60Hz, | 16.44 | ,0) |  |  |  |  | VoltSe (HeatFlux(9),1,MV7_5,16,0,0,_60Hz,16.44,0)
4 |  | 10 | Heat Flux | [6 empty] | PulsePort(1,10000)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 10 | ),1,MV7_5, | 16 | ,0,0,_60Hz, | 15.89 | ,0) |  |  |  |  | VoltSe (HeatFlux(10),1,MV7_5,16,0,0,_60Hz,15.89,0)
4 |  | 11 | Heat Flux | [6 empty] | PulsePort(1,10000)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 11 | ),1,MV7_5, | 16 | ,0,0,_60Hz, | 16.06 | ,0) |  |  |  |  | VoltSe (HeatFlux(11),1,MV7_5,16,0,0,_60Hz,16.06,0)
4 |  | 12 | Heat Flux | [6 empty] | PulsePort(1,10000)
'HFP01 Heat Flux Plates
'Measure heat flux. VoltSe (Dest, Reps, Range, SEChan, MeasOfs, SettlingTime, Integ, Mult, Offset)
VoltSe (HeatFlux( | 12 | ),1,MV7_5, | 16 | ,0,0,_60Hz, | 16.87 | ,0) |  |  |  |  | VoltSe (HeatFlux(12),1,MV7_5,16,0,0,_60Hz,16.87,0)
'Turn AM16/32 Multiplexer Off. PortSet(DataLogger channel C#, State 0=off and reset)
PortSet( | 5 | ,0) | [6 empty] | PortSet(5,0)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MuxNumber | 5 | Linear PAR sensors and Thermocouples | [6 empty] | 'Turn AM16/32 Multiplexer On. PortSet(DataLogger com channel, State 1=on and can respond to clock pulses)
PortSet( | 6 | ,1) | [6 empty] | PortSet(6,1)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 1 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 1 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(1),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 2 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 2 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(2),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 3 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 3 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(3),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 4 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 4 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(4),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 5 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 5 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(5),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 6 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 6 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(6),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 7 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 7 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(7),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 8 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 8 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(8),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 9 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 9 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(9),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 10 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 10 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(10),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 11 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 11 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(11),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 12 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 12 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(12),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 13 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 13 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(13),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 14 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 14 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(14),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 15 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 15 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(15),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the first AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 16 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 16 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(16),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 17 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 17 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(17),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 18 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 18 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(18),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 19 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 19 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(19),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 20 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 20 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(20),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 21 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 21 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(21),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 22 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 22 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(22),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 23 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 23 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(23),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 24 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 24 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(24),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 25 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 25 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(25),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 26 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 26 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(26),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 27 | Linear PAR | [6 empty] | PulsePort(1,10000)
'Measure Voltage Differential for Linear PAR sensor VoltDiff(Dest, Reps, Range, DiffChan, RevDiff, SettlingTime, Integ, Mult, Offset)
VoltDiff(LinPAR( | 27 | ),1,mV2500, | 7 | ,True,0,_60Hz,33.3,0) | [6 empty] | VoltDiff(LinPAR(27),1,mV2500,7,True,0,_60Hz,33.3,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on the next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 9 | Thermocouple | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 9 | ),1,MV2_5C, | 7 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(9),1,MV2_5C,7,TypeE,PTemp_C,True,0,_60Hz,1,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 10 | Thermocouple | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 10 | ),1,MV2_5C, | 7 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(10),1,MV2_5C,7,TypeE,PTemp_C,True,0,_60Hz,1,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 11 | Thermocouple | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 11 | ),1,MV2_5C, | 7 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(11),1,MV2_5C,7,TypeE,PTemp_C,True,0,_60Hz,1,0)
MuxNumber | MUX Chanel (2X32 port) | SensorRef (from SensorCoeff table) | [7 empty] | 'Switch on next AM416 Multiplexer channel. PulsePort(DataLogger channel C#, Delay)
5 |  | 12 | Thermocouple | [6 empty] | PulsePort(1,10000)
'Type E (chromel-constantan) Thermocouple measurements
'Measure average temperature above heat flux plate. TCDiff(Dest, Reps, Range, DiffChan, TCType, Tref, RevDiff, Settlingtime, Integ, Mult, Offset)
TCDiff(SoilAvTC( | 12 | ),1,MV2_5C, | 7 | ,TypeE,PTemp_C,True,0,_60Hz,1,0) | [6 empty] | TCDiff(SoilAvTC(12),1,MV2_5C,7,TypeE,PTemp_C,True,0,_60Hz,1,0)
'Turn AM16/32 Multiplexer Off. PortSet(DataLogger channel C#, State 0=off and reset)
PortSet( | 6 | ,0) | [6 empty] | PortSet(6,0)
Delay(0,150,mSec) 'Delay(Option, Delay, Units)
'Call Data Tables and Store Data
CallTable(AllData)
NextScan
EndProg

## Sheet1
Mux1 | 4x16 | Diff | 16 | IR signal |  |  |  |  | 16
|  | SE excitation | 16 | IR casing |  |  |  |  | 16
Mux2 | 4x16 | Diff | 8 | IR signal | 8 | Thermocouples |  |  | 16
|  | SE excitation | 8 | IR casing | 8 | Heat flux |  |  | 16
Mux3 | 2x32 | Diff | 4 | Thermocouples | 15 | RS intercepted | 12 | SR reflected | 31
Mux4 | 2x32 | SE excitation | 24 | Therm 108 | 4 | Heat flux |  |  | 28
|  |  | IR signal | 24
|  |  | IR casing | 24
|  |  | Thermocouples | 12
|  |  | Heat flux | 12
|  |  | Therm 108 | 24
|  |  | SR reflected | 20
|  |  | RS intercepted | 15